Single Quad Demo

This demo presents a simulation of a single quadrupole mass analyser instrument with collisional cooling interface, with Brubaker prefilter and with account for fringing fields. Note that for this simulations fields are presented by axial distributions obtained by analysis of externally calculated fields with the use of PAnalyser application. Because of this memory requirements for this simulation is below 10MB while presentation of fields is very accurate. For example 2D fields, that presents middle section of analytical quadrupole Q1 is accurate up to the electrode surface. This can be seen on the figure below

Such quality of the field is very important for accurate simulations of quadrupoles and it is difficult to obtain by other field presentations.

Figure below presents main parts of the model and colour gradient shows distribution of pressure in the device. Note that cooling quadrupole features a gas flow of 200m/s. Main parts are obvious from this figure. Analytical quadrupole Q1 has hyperbolic rods with inscribed radius 5mm and total length 200mm. Main operating RF frequency is 2MHz.

Example 1. Optimisation of offset voltages for best transmission.

- Start SIMAX. Using button  load a model SingleQuad_GF200_Opt_609Da_R1k.wb8 from simulation folder.

- Go to SCAN 2D page of the main page control. Press  button.

Software will repeat simulation again and again with different values of Q1_Offset and ST1_Offset voltages as defined by Scan2D page settings. Number of ions detected at the BLOCK region are counted and presented in the plot. Analytical quadrupole is set up to transmit 609Da ions with resolving power 1000. As a result of this simulation the following figure should be obtained:

The best transmission 100% is obtained at ST1 offset=-11V and Q1 offset =-5V.

Example 2. Peak Shape of transmission.

In this simulation a SIMAX tool Scan1D is used. It allows to repeat simulation while changing ion mass or amplitude of some signal in the range and with the step as defined by user. The tool counts splat or flying ions at particular region and presents results as a 1D plot. This tool is very usefull in peak shape simulations or other simulations where transmission of the device is investigated. For present simulation the quadrupole mass analyser Q1 is set up to transmit 609Da ions with resolving power 1000 (V= 4398.0279Vo-p, U= 737.6248V). Based on previous simulation the following signature of offsets is used: IQ0/Q0/IQ1/ST1/Q1/Exit = 10V/0V/-5V/-11V/-5V/-50V

- Start SIMAX. Using button  load a model SingleQuad_GF200_609Da_PeakShape_R1k.wb8 from simulation folder.

- Go to SCAN 1D page of the main page control. Press  button.

As a result of this simulation the following figure should be obtained:

Example 3. Illustration of resolving power 3000.

Here the analytical quadrupole Q1 is set up for resolving power 3000 (V= 4398.0279Vo-p, U= 737.9958). Three groups of ions (100 particles in each group).

Ions have close masses near 609Da and in order to separate those ions one needs resolving power of 3000. Simulation will show this result.

- Start SIMAX. Using button  load a model SingleQuad_609triplet_GF200_R3k.wb8 from simulation folder.

- Select view region by pressing left mouse key at the left-upper point of a WorkBench view and releasing it at the bottom-right corner. Software will repaint region of interest.

- Start ions Fly by pressing  button. Wait until simulation time (10ms) will be completed

As a result of this simulation the following figure should be obtained

At the end of the quadrupole the beam becomes red, which means that all green ions (609.2Da) and blue ions (608.8Da) are rejected to the rods. Figure below shows that transmission of 609Da ions is 93% and some small amount of green ions is also transmitted.